## Calculator
| ASSR(MT) Masking Noise Calculator | [5 empty] | Only for use with modulated tone stimuli |  | V3 2024
Instructions:
Click in pale yellow cells to enter options and appropriate levels then press Enter / Return | [9 empty] | 0 | 0
For single frequency tests, select "Various" and at non-test frequencies, enter 0dB
Warning: some variables will be unknown (e.g. air-bone gaps) so enter several likely values; if in doubt use the higher noise level
This calculator is a time-saving guide; the user must carry clinical responsibility for any decisions they make | [9 empty] | Label | Freq | 1k | -3 to -1
Noise Level is the suggested dB SPL level of noise for the stimulus type, transducer, corrected age & stimulus level | [9 empty] | Click | Wband |  | 0 to 2
| ASSR Equipment | GSI Audera |  |  | Message | Warning! Exclude 500Hz in BC tests |  |  |  | 4k tone | 4k | Include 500 | 3 to 5
Stimulus levels across frequencies: |  | Various |  |  | Area | 100 | moderate |  |  | 2k tone | 2k | Exclude 500 | 6 to 8
Stimulus Transducer: |  | Insert |  |  |  | Warning! There is a moderate risk of cross-masking |  |  |  | 1k tone | 1k |  | 9 to 12
Noise Transducer: |  | Insert | [7 empty] | 500 tone | 500 |  | 13 to 16
Include or exclude 500Hz?: |  | Exclude 500 |  | 1k tone |  | Warning! High Noise Level |  |  |  | CE-Chirp | WBand |  | 17 to 24
Patient corrected age (weeks): |  | 6 to 8 |  |  |  | Warning! Stimulus level exceeds recommended maximum |  |  |  | 4k CE-Chirp | 4k |  | 25 to 104
Test ear air-bone gap, dB: |  | 0 |  | 0 | [5 empty] | 2k CE-Chirp | 2k |  | >104
Non-test ear air-bone gap, dB: |  | 0 | [7 empty] | 1k CE-Chirp | 1k
dBeHL non-test BC threshold: |  | 0 | [5 empty] | -14 | -45 | 500 CE-Chirp | 500
Stimulus Level, dBnHL: |  | 80 |  | 500Hz | 1kHz | 2kHz | 4kHz | No need to mask
Stimulus levels at each frequency, dBnHL: |  |  |  | 50 | 50 | 50 | 50 | 0 |  |  |  |  | Biologic MASTER
Equation: | Stim(dBnHL) | +dBage | +RML | +ABGnt | -IA | -Nt | -Nc | Noise Level |  |  |  |  | GSI Audera
dBnoise = | 80 | 5 | 27 | 0 | 46 | 5 | -9 |  | Off |  |  |  | NeuroAudio
GN EP200/Chartr
Variables for the current settings (dB), based on 500Hz: | [12 empty] | GSI AuderaPro
|  | RMLlower = | 12 | #VALUE! | [8 empty] | Vivosonic Integrity
Nt = | 5 | RMLupper = | 27 | [9 empty] | Custom
Nc = | -9 | ABGnt = | 0
BC age correction = | 20 | IAs = | 46
Stimulus dBeHL = | 60 | IAn = | 46
Stimulus dBeHLnt = | 14 | IAa = | 0
|  | Max noise SPL crossing to test cochlea = | #VALUE!
Please email comments, further data, errors & omisions to: |  |  |  | admin@eratraining.co.uk
For ERA/ABR courses, peer review training/addreditation & clinical support please visit | [5 empty] | www.eratraining.co.uk/
Revision history:
2020: Adapted from the ABR Noise Calculator. For use only with 40Hz & 90Hz modulated tone ASSR stimuli. Not suitable for chirp stimuli.
The noise level at each frequency is calculated and the highest is suggested. Noise is calibrated in dB SPL (BioLogic in HL).
The corrections in Table 8 on the Data tab may not be optimum for Adult 90Hz & 40Hz tests. If you have figures you would prefer to use, enter them in Table 8.
2020 version 2: Maximum stimulus level limited to 100dB to avoid potential noise risk or 80dB at multiple frequencies.
2024 version 3. Updated minimum IA values in table 4 (data tab) from Kariv et al 2024.
Disclaimer: this utility has produced to aid testers but the author can accept no responsibility for errors or inappropriate use.
This is not a medical device; testers should use their own clinical judgement when performing tests on patients.
Note on A9: When testing at a single frequency, select Various then enter stimulus levels at other frequencies to or 0dB to ensure they play no part in the calculated noise level.
Note on A12: BC tests at 500Hz are not advised. In AC tests, 500Hz often takes longer; you may opt to exclude 500Hz and test it separately.
Note on A14: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking. This matters only when testing by air conduction.
Note on A15: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking.
Note on A16: Needed for deciding if masking is needed. This is the best non-test ear BC threshold across frequencies. Enter the lowest value it could be. If unknown select 0.
Note on A17: Enter the highest (across frequencies) stimulus level you intend using - the calculator will then work out whether masking is needed and if so, the level of noise to use. Limit multi-frequency tests to 80dB
Note on A18: Enter the stimulus levels you intend using at each frequency - the calculator will then work out whether masking is needed and if so, the level of noise to use. When testing at a single frequency, enter stimulus levels at other frequencies to 0dB to ensure they play no part in the calculated noise level. When testing at multiple frequencies, limit tests to 80dB.

## 4k (hidden)
ASSR(MT) Masking Noise Calculator - 4kHz
0
[10 empty] | FALSE |  |  | <-3
Result is the dB "dial" level of noise for the selected system, stimulus type, transducer, corrected age & stimulus level | [9 empty] | Label | Freq | 4k | -3 to -1
Offset is an alternative way of specifying noise level: it is the dB noise level relative to the stimulus level | [9 empty] | Click | Wband |  | 0 to 2
[6 empty] | Warning! High Noise Level |  |  |  | 4k tone | 4k |  | 3 to 5
ABR System: |  | GSI Audera |  |  |  | 95 | 0 |  |  | 2k tone | 2k |  | 6 to 8
Stimulus Transducer: |  | Insert |  |  |  | Warning! There is a 0 risk of cross-masking |  |  |  | 1k tone | 1k |  | 9 to 12
Noise Transducer: |  | Insert | [7 empty] | 500 tone | 500 |  | 13 to 16
Stimulus type: |  | 4k tone |  |  |  | Warning! Stimulus level exceeds recommended maximum: |  |  |  | CE-Chirp | WBand |  | 17 to 24
Patient corrected age (weeks): |  | 6 to 8 | [7 empty] | 4k CE-Chirp | 4k |  | 25 to 104
Test ear air-bone gap, dB: |  | 0 |  | 0 | [5 empty] | 2k CE-Chirp | 2k |  | >104
Non-test ear air-bone gap, dB: |  | 0 |  |  |  | 95 |  | dBnHL |  | 1k CE-Chirp | 1k
dBeHL non-test BC ABR threshold: |  | 0 | [5 empty] | -37 | -113 | 500 CE-Chirp | 500
dBnHL Stimulus Level: |  | 50 |  | 35 |  | dBeHL |  | No need to mask
[8 empty] | 0 |  |  |  |  | Biologic MASTER
Equation: | Stim(dBnHL) | +dBage | +RML | +ABGnt | -IA | -Nt | -Nc | Result | Offset |  |  |  | Interacoustics Eclipse
dBnoise = | 50 | 10 | 25 | 0 | 74 | 10 | -9 | 10 | -40 |  |  |  | GSI Audera
NeuroAudio
Variables for the current settings (dB): | [12 empty] | GN EP200/Chartr
|  | RMLlower = | 10 | 11 | [8 empty] | GSI AuderaPro
Nt = | 10 | RMLupper = | 25 | [9 empty] | Vivosonic Integrity
Nc = | -9 | ABGnt = | 0 | [9 empty] | Custom
BC age correction = | 10 | IAs = | 54
Stimulus dBeHL = | 35 | IAn = | 54
Stimulus dBeHLnt = | -39 | IAa = | 20
|  | Max noise SPL crossing to test cochlea = | -63
Disclaimer: the author has produced this in good faith but can accept no responsibility for errors or inappropriate use.
Note on A14: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking. This matters only when testing by air conduction.
Note on A15: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking.
Note on A16: Needed for deciding if masking is needed. This is the non-test ear BC threshold for the chosen stimulus type. Enter the lowest value it could be. If unknown select 0.
Note on A17: Enter the test level you are using - the calculator will then work out whether masking is needed and if so, the level of noise to use.
Note on C19: Age-related correction added to stimulus level. This accounts for BC or insert age effects.
Note on D19: RMLupper: The level of noise guaranteed to mask the stimulus in the same ear. Assumes stimulus level in dBnHL and noise level in dBSPL
Note on E19: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked
Note on F19: Minimum Interaural Attenuation (transcanial transmission loss) This depends on stimulus and transducer. Plus: Additional age-related IA (TTL) (incomplete skull fusion) applicable to all transducers.
Note on G19: Additional age-related correction when inserts are used as the noise transducer. Depends on age.
Note on H19: Nc is the noise calibration offset for the selected system and noise transducer. This is set by the manufacturer - there is no international standard for white noise.
Note on I19: The result has been rounded to the nearest 5dB step
Note on J19: Some systems have the option of setting the noise level as an "offset" - this is handy since changes in the stimulus level automatically change the noise level (synchronous masking)
Note on C23: Level of noise that will just not totally mask the stimulus in the same ear of any subject. Used in assessing risk of cross-masking. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A24: Age-related insert transducer SPL lift correction for age. Applies to stimulus & noise when applied via inserts. Data Table 5. Depends on age & stimulus.
Note on C24: Level of noise guaranteed to mask the stimulus in the same ear in all subjects. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A25: Nc is the noise calibration offset for the selected system. Data Table 1. Depends on selected noise transducer
Note on C25: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked.
Note on A26: BC lift caused by incomplete cranial bone fusion. Data Table 6. Click data based on Webb (1993). This value is included in the equation only when BC is used.
Note on C26: Minimum interaural attenuation (transcanial transmission loss) for the stimulus. Data Table 3. Depends on transducer and stimulus.
Note on A27: For information only. Applies NHSP corrections for click BC/Insert age effects and nHL-eHL offsets.
Note on C27: Minimum interaural attenuation (transcanial transmission loss) for the noise. Data Table 3. Depends on transducer and stimulus.
Note on A28: Maximum stimulus level in dBeHL reaching non-test cochlea. This is helpful when deciding whether masking is necessary.
Note on C28: Additional age-related IA (TTL). Applicable to all transducers. Data Table 4: IA depends on age & stimulus
Note on C29: Used in calculating the risk of cross-masking. Maximum noise level at the test ear cochlea in dB SPL

## 2k (hidden)
ASSR(MT) Masking Noise Calculator - 2kHz
0
[10 empty] | FALSE |  |  | <-3
Result is the dB "dial" level of noise for the selected system, stimulus type, transducer, corrected age & stimulus level | [9 empty] | Label | Freq | 2k | -3 to -1
Offset is an alternative way of specifying noise level: it is the dB noise level relative to the stimulus level | [9 empty] | Click | Wband |  | 0 to 2
[6 empty] | Warning! High Noise Level |  |  |  | 4k tone | 4k |  | 3 to 5
ABR System: |  | GSI Audera |  |  |  | 100 | 0 |  |  | 2k tone | 2k |  | 6 to 8
Stimulus Transducer: |  | Insert |  |  |  | Warning! There is a 0 risk of cross-masking |  |  |  | 1k tone | 1k |  | 9 to 12
Noise Transducer: |  | Insert | [7 empty] | 500 tone | 500 |  | 13 to 16
Stimulus type: |  | 2k tone |  |  |  | Warning! Stimulus level exceeds recommended maximum: |  |  |  | CE-Chirp | WBand |  | 17 to 24
Patient corrected age (weeks): |  | 6 to 8 | [7 empty] | 4k CE-Chirp | 4k |  | 25 to 104
Test ear air-bone gap, dB: |  | 0 |  | 0 | [5 empty] | 2k CE-Chirp | 2k |  | >104
Non-test ear air-bone gap, dB: |  | 0 |  |  |  | 100 |  | dBnHL |  | 1k CE-Chirp | 1k
dBeHL non-test BC ABR threshold: |  | 0 | [5 empty] | -17 | -69 | 500 CE-Chirp | 500
dBnHL Stimulus Level: |  | 50 |  | 35 |  | dBeHL |  | No need to mask
[8 empty] | 0 |  |  |  |  | Biologic MASTER
Equation: | Stim(dBnHL) | +dBage | +RML | +ABGnt | -IA | -Nt | -Nc | Result | Offset |  |  |  | Interacoustics Eclipse
dBnoise = | 50 | 5 | 26 | 0 | 54 | 5 | -9 | 30 | -20 |  |  |  | GSI Audera
NeuroAudio
Variables for the current settings (dB): | [12 empty] | GN EP200/Chartr
|  | RMLlower = | 6 | 26 | [8 empty] | GSI AuderaPro
Nt = | 5 | RMLupper = | 26 | [9 empty] | Vivosonic Integrity
Nc = | -9 | ABGnt = | 0 | [9 empty] | Custom
BC age correction = | 5 | IAs = | 44
Stimulus dBeHL = | 35 | IAn = | 44
Stimulus dBeHLnt = | -19 | IAa = | 10
|  | Max noise SPL crossing to test cochlea = | -28
Disclaimer: the author has produced this in good faith but can accept no responsibility for errors or inappropriate use.
Note on A14: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking. This matters only when testing by air conduction.
Note on A15: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking.
Note on A16: Needed for deciding if masking is needed. This is the non-test ear BC threshold for the chosen stimulus type. Enter the lowest value it could be. If unknown select 0.
Note on A17: Enter the test level you are using - the calculator will then work out whether masking is needed and if so, the level of noise to use.
Note on C19: Age-related correction added to stimulus level. This accounts for BC or insert age effects.
Note on D19: RMLupper: The level of noise guaranteed to mask the stimulus in the same ear. Assumes stimulus level in dBnHL and noise level in dBSPL
Note on E19: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked
Note on F19: Minimum Interaural Attenuation (transcanial transmission loss) This depends on stimulus and transducer. Plus: Additional age-related IA (TTL) (incomplete skull fusion) applicable to all transducers.
Note on G19: Additional age-related correction when inserts are used as the noise transducer. Depends on age.
Note on H19: Nc is the noise calibration offset for the selected system and noise transducer. This is set by the manufacturer - there is no international standard for white noise.
Note on I19: The result has been rounded to the nearest 5dB step
Note on J19: Some systems have the option of setting the noise level as an "offset" - this is handy since changes in the stimulus level automatically change the noise level (synchronous masking)
Note on C23: Level of noise that will just not totally mask the stimulus in the same ear of any subject. Used in assessing risk of cross-masking. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A24: Age-related insert transducer SPL lift correction for age. Applies to stimulus & noise when applied via inserts. Data Table 5. Depends on age & stimulus.
Note on C24: Level of noise guaranteed to mask the stimulus in the same ear in all subjects. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A25: Nc is the noise calibration offset for the selected system. Data Table 1. Depends on selected noise transducer
Note on C25: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked.
Note on A26: BC lift caused by incomplete cranial bone fusion. Data Table 6. Click data based on Webb (1993). This value is included in the equation only when BC is used.
Note on C26: Minimum interaural attenuation (transcanial transmission loss) for the stimulus. Data Table 3. Depends on transducer and stimulus.
Note on A27: For information only. Applies NHSP corrections for click BC/Insert age effects and nHL-eHL offsets.
Note on C27: Minimum interaural attenuation (transcanial transmission loss) for the noise. Data Table 3. Depends on transducer and stimulus.
Note on A28: Maximum stimulus level in dBeHL reaching non-test cochlea. This is helpful when deciding whether masking is necessary.
Note on C28: Additional age-related IA (TTL). Applicable to all transducers. Data Table 4: IA depends on age & stimulus
Note on C29: Used in calculating the risk of cross-masking. Maximum noise level at the test ear cochlea in dB SPL

## 1k (hidden)
ASSR(MT) Masking Noise Calculator - 1kHz
0
[10 empty] | FALSE |  |  | <-3
Result is the dB "dial" level of noise for the selected system, stimulus type, transducer, corrected age & stimulus level | [9 empty] | Label | Freq | 1k | -3 to -1
Offset is an alternative way of specifying noise level: it is the dB noise level relative to the stimulus level | [9 empty] | Click | Wband |  | 0 to 2
[6 empty] | Warning! High Noise Level |  |  |  | 4k tone | 4k |  | 3 to 5
ABR System: |  | GSI Audera |  |  |  | 100 | 0 |  |  | 2k tone | 2k |  | 6 to 8
Stimulus Transducer: |  | Insert |  |  |  | Warning! There is a 0 risk of cross-masking |  |  |  | 1k tone | 1k |  | 9 to 12
Noise Transducer: |  | Insert | [7 empty] | 500 tone | 500 |  | 13 to 16
Stimulus type: |  | 1k tone |  |  |  | Warning! Stimulus level exceeds recommended maximum: |  |  |  | CE-Chirp | WBand |  | 17 to 24
Patient corrected age (weeks): |  | 6 to 8 | [7 empty] | 4k CE-Chirp | 4k |  | 25 to 104
Test ear air-bone gap, dB: |  | 0 |  | 0 | [5 empty] | 2k CE-Chirp | 2k |  | >104
Non-test ear air-bone gap, dB: |  | 0 |  |  |  | 100 |  | dBnHL |  | 1k CE-Chirp | 1k
dBeHL non-test BC ABR threshold: |  | 0 | [5 empty] | -14 | -57 | 500 CE-Chirp | 500
dBnHL Stimulus Level: |  | 50 |  | 30 |  | dBeHL |  | No need to mask
[8 empty] | 0 |  |  |  |  | Biologic MASTER
Equation: | Stim(dBnHL) | +dBage | +RML | +ABGnt | -IA | -Nt | -Nc | Result | Offset |  |  |  | Interacoustics Eclipse
dBnoise = | 50 | 5 | 27 | 0 | 46 | 5 | -9 | 40 | -10 |  |  |  | GSI Audera
NeuroAudio
Variables for the current settings (dB): | [12 empty] | GN EP200/Chartr
|  | RMLlower = | 12 | 36 | [8 empty] | GSI AuderaPro
Nt = | 5 | RMLupper = | 27 | [9 empty] | Vivosonic Integrity
Nc = | -9 | ABGnt = | 0 | [9 empty] | Custom
BC age correction = | 20 | IAs = | 46
Stimulus dBeHL = | 30 | IAn = | 46
Stimulus dBeHLnt = | -16 | IAa = | 0
|  | Max noise SPL crossing to test cochlea = | -10
Disclaimer: the author has produced this in good faith but can accept no responsibility for errors or inappropriate use.
Note on A14: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking. This matters only when testing by air conduction.
Note on A15: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking.
Note on A16: Needed for deciding if masking is needed. This is the non-test ear BC threshold for the chosen stimulus type. Enter the lowest value it could be. If unknown select 0.
Note on A17: Enter the test level you are using - the calculator will then work out whether masking is needed and if so, the level of noise to use.
Note on C19: Age-related correction added to stimulus level. This accounts for BC or insert age effects.
Note on D19: RMLupper: The level of noise guaranteed to mask the stimulus in the same ear. Assumes stimulus level in dBnHL and noise level in dBSPL
Note on E19: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked
Note on F19: Minimum Interaural Attenuation (transcanial transmission loss) This depends on stimulus and transducer. Plus: Additional age-related IA (TTL) (incomplete skull fusion) applicable to all transducers.
Note on G19: Additional age-related correction when inserts are used as the noise transducer. Depends on age.
Note on H19: Nc is the noise calibration offset for the selected system and noise transducer. This is set by the manufacturer - there is no international standard for white noise.
Note on I19: The result has been rounded to the nearest 5dB step
Note on J19: Some systems have the option of setting the noise level as an "offset" - this is handy since changes in the stimulus level automatically change the noise level (synchronous masking)
Note on C23: Level of noise that will just not totally mask the stimulus in the same ear of any subject. Used in assessing risk of cross-masking. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A24: Age-related insert transducer SPL lift correction for age. Applies to stimulus & noise when applied via inserts. Data Table 5. Depends on age & stimulus.
Note on C24: Level of noise guaranteed to mask the stimulus in the same ear in all subjects. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A25: Nc is the noise calibration offset for the selected system. Data Table 1. Depends on selected noise transducer
Note on C25: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked.
Note on A26: BC lift caused by incomplete cranial bone fusion. Data Table 6. Click data based on Webb (1993). This value is included in the equation only when BC is used.
Note on C26: Minimum interaural attenuation (transcanial transmission loss) for the stimulus. Data Table 3. Depends on transducer and stimulus.
Note on A27: For information only. Applies NHSP corrections for click BC/Insert age effects and nHL-eHL offsets.
Note on C27: Minimum interaural attenuation (transcanial transmission loss) for the noise. Data Table 3. Depends on transducer and stimulus.
Note on A28: Maximum stimulus level in dBeHL reaching non-test cochlea. This is helpful when deciding whether masking is necessary.
Note on C28: Additional age-related IA (TTL). Applicable to all transducers. Data Table 4: IA depends on age & stimulus
Note on C29: Used in calculating the risk of cross-masking. Maximum noise level at the test ear cochlea in dB SPL

## 500 (hidden)
ASSR(MT) Masking Noise Calculator - 500Hz
0
[10 empty] | FALSE |  |  | <-3
This calculator is a time-saving guide; the user must carry clinical responsibility for any decisions they make
Result is the dB "dial" level of noise for the selected system, stimulus type, transducer, corrected age & stimulus level | [9 empty] | Label | Freq | 500 | -3 to -1
Offset is an alternative way of specifying noise level: it is the dB noise level relative to the stimulus level | [9 empty] | Click | Wband |  | 0 to 2
[6 empty] | Warning! High Noise Level |  |  |  | 4k tone | 4k |  | 3 to 5
ABR System: |  | GSI Audera |  |  |  | 100 | 0 |  |  | 2k tone | 2k |  | 6 to 8
Stimulus Transducer: |  | Insert |  |  |  | Warning! There is a 0 risk of cross-masking |  |  |  | 1k tone | 1k |  | 9 to 12
Noise Transducer: |  | Insert | [7 empty] | 500 tone | 500 |  | 13 to 16
Stimulus type: |  | 500 tone |  |  |  | Warning! Stimulus level exceeds recommended maximum: |  |  |  | CE-Chirp | WBand |  | 17 to 24
Patient corrected age (weeks): |  | 6 to 8 | [7 empty] | 4k CE-Chirp | 4k |  | 25 to 104
Test ear air-bone gap, dB: |  | 0 |  | 0 | [5 empty] | 2k CE-Chirp | 2k |  | >104
Non-test ear air-bone gap, dB: |  | 0 |  |  |  | 100 |  | dBnHL |  | 1k CE-Chirp | 1k
dBeHL non-test BC ABR threshold: |  | 0 | [5 empty] | -88 | -45 | 500 CE-Chirp | 500
dBnHL Stimulus Level: |  | -30 |  | -50 |  | dBeHL |  | No need to mask
[8 empty] | 0 |  |  |  |  | Biologic MASTER
Equation: | Stim(dBnHL) | +dBage | +RML | +ABGnt | -IA | -Nt | -Nc | Result | Offset |  |  |  | Interacoustics Eclipse
dBnoise = | -30 | 5 | 26 | 0 | 40 | 5 | -9 | -35 | -5 |  |  |  | GSI Audera
NeuroAudio
Variables for the current settings (dB): | [12 empty] | GN EP200/Chartr
|  | RMLlower = | 11 | -39 | [8 empty] | GSI AuderaPro
Nt = | 5 | RMLupper = | 26 | [9 empty] | Vivosonic Integrity
Nc = | -9 | ABGnt = | 0 | [9 empty] | Custom
BC age correction = | 25 | IAs = | 40
Stimulus dBeHL = | -50 | IAn = | 40
Stimulus dBeHLnt = | -90 | IAa = | 0
|  | Max noise SPL crossing to test cochlea = | -79
Disclaimer: the author has produced this in good faith but can accept no responsibility for errors or inappropriate use.
Note on A14: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking. This matters only when testing by air conduction.
Note on A15: This is normally unknown so estimate, on the basis of tymp & any other available clinical information. It influences the risk of cross-masking: the greater the air-bone gap, the greater the likelihood of cross-masking.
Note on A16: Needed for deciding if masking is needed. This is the non-test ear BC threshold for the chosen stimulus type. Enter the lowest value it could be. If unknown select 0.
Note on A17: Enter the test level you are using - the calculator will then work out whether masking is needed and if so, the level of noise to use.
Note on C19: Age-related correction added to stimulus level. This accounts for BC or insert age effects.
Note on D19: RMLupper: The level of noise guaranteed to mask the stimulus in the same ear. Assumes stimulus level in dBnHL and noise level in dBSPL
Note on E19: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked
Note on F19: Minimum Interaural Attenuation (transcanial transmission loss) This depends on stimulus and transducer. Plus: Additional age-related IA (TTL) (incomplete skull fusion) applicable to all transducers.
Note on G19: Additional age-related correction when inserts are used as the noise transducer. Depends on age.
Note on H19: Nc is the noise calibration offset for the selected system and noise transducer. This is set by the manufacturer - there is no international standard for white noise.
Note on I19: The result has been rounded to the nearest 5dB step
Note on J19: Some systems have the option of setting the noise level as an "offset" - this is handy since changes in the stimulus level automatically change the noise level (synchronous masking)
Note on C23: Level of noise that will just not totally mask the stimulus in the same ear of any subject. Used in assessing risk of cross-masking. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A24: Age-related insert transducer SPL lift correction for age. Applies to stimulus & noise when applied via inserts. Data Table 5. Depends on age & stimulus.
Note on C24: Level of noise guaranteed to mask the stimulus in the same ear in all subjects. Assumes stimulus in dBnHL and noise in dBSPL. Data Table 2.
Note on A25: Nc is the noise calibration offset for the selected system. Data Table 1. Depends on selected noise transducer
Note on C25: Air-bone gap in the non-test ear. Noise needs to be increased by this amount to overcome the effect of conductive element in the ear being masked.
Note on A26: BC lift caused by incomplete cranial bone fusion. Data Table 6. Click data based on Webb (1993). This value is included in the equation only when BC is used.
Note on C26: Minimum interaural attenuation (transcanial transmission loss) for the stimulus. Data Table 3. Depends on transducer and stimulus.
Note on A27: For information only. Applies NHSP corrections for click BC/Insert age effects and nHL-eHL offsets.
Note on C27: Minimum interaural attenuation (transcanial transmission loss) for the noise. Data Table 3. Depends on transducer and stimulus.
Note on A28: Maximum stimulus level in dBeHL reaching non-test cochlea. This is helpful when deciding whether masking is necessary.
Note on C28: Additional age-related IA (TTL). Applicable to all transducers. Data Table 4: IA depends on age & stimulus
Note on C29: Used in calculating the risk of cross-masking. Maximum noise level at the test ear cochlea in dB SPL

## Notes
Notes:
1. ASSR corrections, BC & Insert age corrections taken from BSA 2020 Early Assessment Guidelines.
2. In ASSR tests, when using stimulus levels above 80dB always revert to single frequency testing.
3. Assumptions: ASSR stimuli are calibrated to ISO 389-1/2/3 (pure tones) without modulation.
4. Warning of stimulus level exceeding BSA recommended maximum appears for all ages except >24 weeks.
5. Masking is needed if the stimulus level reaching the non-test cochlea exceeds the non-test BC ABR threshold (both in dBeHL).
6. Cross-masking is considered a posibility if the noise SPL reaching the test ear cochlea exceeds the corrected stimulus level (in dBnHL) reaching the test ear cochlea plus the lower EML minus 20dB (moderate risk) or 10dB (high risk)
7. The corrections in Table 8 on the Data tab may not be optimum for Adult 90Hz & 40Hz tests. If you have figures you would prefer to use, enter them in Table 8.
See also notes on Data Sheet
Don't have Microsoft Excel on your ABR system? (you must be using another computer at the moment, otherwise you wouldn't be reading this!)
You can download a free Excel viewer (you may need to ask your friendly local IT department to do this for you) for computers or as a smartphone app

## Data
Table 1: ABR system noise calibration. dB SPL noise for 0dBdial |  |  |  | Table 2: RML values Assumes noise calibrated in SPL |  |  | Table 3: Minimum adult IAs versus transducer & stimulus |  |  |  | Table 4: IAa. Additional age-related minimum IA versus stimulus | [5 empty] | Table 5: Insert canal volume SPL lift (used for both stimulus and noise) | [5 empty] | Table 6. BC age correction. | [5 empty] | Table 7. Maximum AC stimulus levels |  |  | Table 8. nHL-eHL offsets
System | Nc Phones | Nc Ins/2cc | Nc Ins/711 | Stimulus | RMLupper | RMLlower | Stimulus | IA phones | IA inserts | IA BC | Corrected age (weeks) | Wband | 4k | 2k | 1k | 500 | Corrected age (weeks) | Wband | 4k | 2k | 1k | 500 | Corrected age (weeks) | Wband | 4k | 2k | 1k | 500 | Stimulus | Phones | Inserts | Stimulus | dB
Biologic MASTER | 21 | 19 | 28 | Click | 33 | 18 | Click | 38 | 45 | 0 | <-3 | 20 | 20 | 10 | 0 | 0 | <-3 | 10 | 10 | 5 | 5 | 5 | <-3 | 12 | 10 | 5 | 20 | 25 | Click | 100 | 85 | Click | 5
GSI Audera | 0 | -9 | 0 | 4k tone | 25 | 10 | 4k tone | 42 | 54 | 0 | -3 to -1 | 20 | 20 | 10 | 0 | 0 | -3 to -1 | 10 | 10 | 5 | 5 | 5 | -3 to -1 | 10 | 10 | 5 | 20 | 25 | 4k tone | 115 | 95 | 4k tone | 25
NeuroAudio | 0 | -9 | 0 | 2k tone | 26 | 6 | 2k tone | 35 | 44 | 0 | 0 to 2 | 20 | 20 | 10 | 0 | 0 | 0 to 2 | 10 | 10 | 5 | 5 | 5 | 0 to 2 | 8 | 10 | 5 | 20 | 25 | 2k tone | 120 | 100 | 2k tone | 20
GN EP200/Chartr | 0 | 0 | 9 | 1k tone | 27 | 12 | 1k tone | 37 | 46 | 0 | 3 to 5 | 20 | 20 | 10 | 0 | 0 | 3 to 5 | 10 | 10 | 5 | 5 | 5 | 3 to 5 | 6 | 10 | 5 | 20 | 25 | 1k tone | 120 | 100 | 1k tone | 25
GSI AuderaPro | 0 | -9 | 0 | 500 tone | 26 | 11 | 500 tone | 35 | 40 | 0 | 6 to 8 | 20 | 20 | 10 | 0 | 0 | 6 to 8 | 10 | 10 | 5 | 5 | 5 | 6 to 8 | 5 | 10 | 5 | 20 | 25 | 500 tone | 120 | 100 | 500 tone | 25
Vivosonic Integrity | 0 | -9 | 0 | CE-Chirp | 33 | 18 | CE-Chirp | 38 | 45 | 0 | 9 to 12 | 20 | 20 | 10 | 0 | 0 | 9 to 12 | 10 | 10 | 5 | 5 | 5 | 9 to 12 | 3 | 10 | 5 | 20 | 25 | CE-Chirp | 100 | 85 | CE-Chirp | 0
|  |  |  | 4k CE-Chirp | 28 | 13 | 4k CE-Chirp | 42 | 54 | 0 | 13 to 16 | 15 | 15 | 0 | 0 | 0 | 13 to 16 | 5 | 5 | 0 | 0 | 0 | 13 to 16 | 0 | 5 | 0 | 15 | 20 | 4k CE-Chirp | 105 | 85 | 4k CE-Chirp | 5
Custom |  |  |  | 2k CE-Chirp | 33 | 13 | 2k CE-Chirp | 35 | 44 | 0 | 17 to 24 | 5 | 15 | 0 | 0 | 0 | 17 to 24 | 5 | 5 | 0 | 0 | 0 | 17 to 24 | 0 | 5 | 0 | 15 | 20 | 2k CE-Chirp | 110 | 95 | 2k CE-Chirp | 5
Notes: Based on a sample of available systems (6cc NBS 9A and 2cc HA-2 couplers) or manufacturer information and assumes the noise calibration of your system is similar - check with your calibration agent. Ensure any "offset" in the test protocol is zero when calibrating the noise level and that the transducer type being used matches the type for the calibration file. Custom values may be entered if desired. |  |  |  | 1k CE-Chirp | 28 | 13 | 1k CE-Chirp | 37 | 46 | 0 | 25 to 104 | 0 | 0 | 0 | 0 | 0 | 25 to 104 | 0 | 0 | 0 | 0 | 0 | 25 to 104 | 0 | 0 | 0 | 10 | 15 | 1k CE-Chirp | 115 | 100 | 1k CE-Chirp | 10
|  |  |  | 500 CE-Chirp | 33 | 18 | 500 CE-Chirp | 35 | 40 | 0 | >104 | 0 | 0 | 0 | 0 | 0 | >104 | 0 | 0 | 0 | 0 | 0 | >104 | 0 | 0 | 0 | 0 | 0 | 500 CE-Chirp | 110 | 100 | 500 CE-Chirp | 15
|  |  |  | Noise in each third-octave band will be 13-17dB below WB level (Lightfoot 2010, Fig 1). For RML upper, add + 5dB for effectiveness plus the values given in ISO 389-4. RML lower are these minus the upper-lower range given in Lightfoot et all 2010 |  |  | From ABR noise calculator. Notes: Data on how IA changes with frequency is incomplete. Further research is needed. Source: TDH: average of a number of published studies for pure tones, lower 95% confidence level. Inserts: Munro & Contractor, 2010, minimum of range in 15 subjects. BC: 0dB taken as conservative value. 2012: TDH & insert IA values have been reduced by 10dB at all frequencies to address concerns over undermasking by AC. |  |  |  | Notes: IA is assumed to be greater in immature skulls because of incomplete fusion of cranial bones. Further research is needed to define how these values change with age and frequency. Tonal data up to 1 year from Kariv et al 2024. | [5 empty] | Notes. From BSA Early Assessment Guidelines, 2020. Further research is needed to further refine these data. | [5 empty] | Notes. 1k & 500Hz: Data from study by Ferm 2013. Applicable to clicks, tones & chirps. | [5 empty] | Notes: Modified from BSA Early Assessment Guidelines, 2020. The spreadsheet applies these levels up to 24 weeks corrected age. The reference levels for tone pip and continuous tone stimuli are at least 10dB apart so the above maximum levels are cautious. |  |  | Offsets vary by system, method and state of patient. If you have your own reference data, enter it here. The BSA default values are: 4k tone: 25dB 2k tone: 20dB 1k tone: 25dB 500 tone: 25dB
All systems use modulated tone stimuli
Vivosonic uses their own chirps but switches to tones at high levels so this calculator applies only to high stim levels, with modulated tones | [31 empty] | From BSA Early Assessment 2020 at low levels (phones) i.e. do not include offsets for inserts or BC (added separately)
Note on B10: Enter your system's noise dBSPL above the dial level, measured in an NBS 9A 6cc coupler
Note on C10: Enter your system's noise dBSPL above the dial level, measured in an HA-2 2cc coupler
Note on D10: Enter your system's noise dBSPL above the dial level, measured in a 711 occluded ear simulator then subtract 9dB and enter that value in the cell to the left